# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title : Supply of Equipment for Strengthening Forensic Capacity of Turkey - Relaunch	p 1 /…

Publication reference : EuropeAid/137473/IH/SUP/TR

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

ANNEX II:  TECHNICAL SPECIFICATIONS

- GENERAL REQUIREMENTS

1.1	General Technical Requirements

1.1.1 		All the equipment shall be provided complete with the necessary accessories and/or parts to ensure that the unit is capable of operating to the required technical and quality specifications immediately. All specifications details listed within lot for each item is the minimum requirements. Any improvements on the specifications or additional features offered shall be clearly identified in the Tenderer’s offer.

1.1.2		The main supply voltage in Turkey is 220 V (single-phase) and 380 V AC (tri-phase) with neutral, 50 Hz nominal. All equipment shall be operated on 220 V ± 20 V, 50 Hz ± 0.5 Hz, or 380 V ± 40 V power supply and be suitable for direct connection to the standard single phase/three phase power outlets in Turkey. In addition to power black-outs, the quality and stability of the supplied voltage may undergo fluctuations (plus or minus) of more than 10%. Therefore, necessary back-up power supplies and voltage stabilizers shall be provided to ensure full availability the equipment.

1.1.3		The supplier will evaluate the supplied current, the quality of the current and the fluctuations of the current and take the necessary precautions to avoid damages to the equipment.

1.1.4		The items provided shall not be hybrids and shall be a registered brand name. Because of the variety of equipment it is impossible to put the “one manufacturer” condition, thus for compatibility purposes it is required to follow the compatibility requirements very strictly (PC compatibility, graphics compatibility etc.).

1.1.5		The equipment must have quality certificate to prove that it complies with European standards. All items supplied conforming to the necessary CE regulation/norm must carry a suitable CE badge of conformity, permanently fixed to the machine, where applicable.

1.1.6		It should be noted that whenever a specific name of a product is mentioned in the Technical Specifications, a sufficiently precise and fully intelligible description is not possible, and it has to be understood as that product or its equivalent. Where equivalency shall be subject to technical evaluation, the respective documentation of equivalency, and - if appropriate - an assessment by an independent party shall be provided with the offer.

1.1.7		Equipment which allows upgrading of capacities (except the power capacity) shall be provided in such a way that upgrades can be performed by installing additional capacity without discarding the already installed capacities.

1.1.8		All equipment shall furthermore comply with the state of the art concerning safe operation, energy saving and environmental safety, in line with the generic policies or the practical measures in use within the EU, namely within the scope of Product Safety, Energy Efficiency and Environment Management.

1.1.9		Whole system and system components and equipment in lot shall be compatible each other.

1.1.10		Supply delivery, including installation, integration and final customization must include all needed parts, accessories and consumables required for the supplies to be presented for provisional acceptance fully installed, operational and ready for use.	Consumables, accessories, parts and documentation used during delivery, installation, integration and customization before provisional acceptance must therefore be anticipated and calculated into the offer.	It shall be the sole responsibility of the Tenderer to ensure that all pre-requisites for the completeness of the supply delivery are met before its commencement.

1.1.11		The Tenderer must keep the work site(s) clean and safe against any hazard (e.g. fire, etc.) during all supply stages until provisional acceptance.

1.1.12		The Tenderer must provide the necessary measures to prevent any damage during delivery, installation, integration and customization stages. If any damage occurs during delivery, installation, integration and customization stages, the Tenderer must appropriately rectify it.

1.1.13		When required by national and/or international regulations, all the costs relating to any permission, licensing, certification, testing, calibration or other documentation of the supplies shall be borne by the Contractor.

1.1.14		For the items using GSM network, if required by national and/or international regulations, the above mentioned procedures will be carried out and all the costs will be borne by the Contractor.

1.1.15		There shall not be any usernames, passwords, services, etc. providing access to the proposed products out of the beneficiary institution’s knowledge.

1.1.16		There shall not be any mechanism on the proposed products that provides access for the third parties to any kind of the institution’s data out of the Beneficiary institution’s knowledge. Otherwise, the contractor will be hold judicial and administratively responsible on a related illegal occasion.

1.2		Software Licensing

1.2.1		All software must be licensed to the Beneficiary allowing designated personnel to perform diagnostics, installation, update/upgrade and repair/debug activities without any external assistance.

1.2.2		For any item containing software offered under commercial licensing model all software must be provided with unlimited duration license(s) and technical support within warranty period.

1.3	General Training Requirements

The Supplier shall provide an adequate training for a number of end users as specified in the below table in a way that they will be able to operate and do the maintenance of the equipment without any help.

The minimum number of end users to be trained and the minimum period of the training is given in detail below.

| ITEM | DESCRIPTION | NUMBER OF TRAINEES |  | DURATION (HOURS) |  |
| --- | --- | --- | --- | --- | --- |
|  |  | OPERATIONAL | MAINTENANCE | OPERATIONAL | MAINTENANCE |
| Lot 1 | Laboratory Analysis Devices |  |  |  |  |
| 1.1 | Scanning Electron Microscope-Energy Dispersive X-Ray Detector (SEM-EDS) | 3 | 3 | 6 | 2 |
| 1.2 | Gas Chromatography-Mass Spectrometry System (GC-MS) | 3 | 3 | 5 | 2 |

Trainings will be carried out at the installation locations. The information for number of tarinees and duration in the above table is for each installation location.

Training will cover calibration, training in instrument operation including operation tasks and also training for user maintenance.

The language of the training and the training documents will be Turkish; if not, interpretation/translation will be provided and paid by the Supplier.

The contractor will be responsible to provide all the necessary service, technical personnel, product, equipment, documentation for the training programs.

The Contractor shall provide easy way for calibration of embedded instruments to ensure their metrological traceability. Also the Supplier must include in his proposal the methodologies for the training.

The supplier shall provide training in use of Software and reporting.

This proposal shall include:

  - The type of training

  - Hours/day offered for each item

  - Qualifications of trainers

  - Description of methodology

The trainings for should be delivered by the certificated trainers (unless a certificate is not existed).

Starting date and time-table of the training courses should be determined by mutual agreement between the beneficiary and the supplier.

1.4		General Requirements for Technical Documents to be delivered during implementation period

1.4.1		Each set of equipment shall be provided with an original set of users’ manuals for all parts of the equipment. These manuals shall be supplied in Turkish or English language. If manuals are not in Turkish, a “Quick Guide” in Turkish shall be supplied together with the manual where applicable.

2. 	VISIBILITY

All supplies must comply with the generic Visibility policies in force within the scope of external aid contracts financed from the EU general budget; tenderers will thus be aware that certain visibility rules apply and that the guidelines and manuals concerned may be found in the EUROPEAID website, at:

http://ec.europa.eu/europeaid/funding/communication-and-visibility-manual-eu-external-actions_en

http://www.avrupa.info.tr/AB_Mali_Destegi/Gorunurluk,Visi.html?LanguageID=2

Unless otherwise specified, all equipment shall have a solidly fixed and durable label, size 75 mm x 35 mm, as appropriate for each equipment, with the standard EU logo as below:

INSTRUMENT FOR PRE-ACCESSION ASSISTANCE (IPA) PROGRAMME 2007-2013

EuropeAid/137473/IH/SUP/TR

Lot No:

Item No:

Serial No:

%75 Avrupa Birliği Katkısı, %25 Ulusal Katkı ile alınmıştır

EU Contribution 75% National Contribution 25%

3. 	TECHNICAL SPECIFICATIONS TABLE

| Item No | Description |
| --- | --- |
| 1.1 | Scanning Electron Microscope-Energy Dispersive X-Ray Detector (SEM-EDS) |
| 1.2 | Gas Chromatography-Mass Spectrometry System (GC-MS) |

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1.1 | Scanning Electron Microscope with Energy Dispersive X-Ray Detector (SEM-EDS) System shall be able to be used for qualitative, quantitative and comparative analysis of gunshot residues, inorganic materials, explosives, paints, coatings, metallic substances, liquids, fibers and tool marks at Forensic Laboratories (Police and Gendarmerie Laboratories). |  |  |  |
| 1.1.0.1 | System consists of: Scanning Electron Microscope (electron gun and optical system; Sample compartment, Sample stage and Sample holders; Vacuum system; Detectors), Energy Dispersive X-ray Spectrometer, System Computer, Software, Sample Coating Unit and accessories. |  |  |  |
| 1.1.0.2 | General Requirements: |  |  |  |
| 1.1.0.2.1 | Systems shall work at high vacuum, low vacuum and environmental mode conditions and shall be switched from one to another and shall automatically adjust itself. |  |  |  |
| 1.1.0.2.2 | Systems shall produce secondary electron (SE) and backscattered electron (BSE) images. |  |  |  |
| 1.1.0.2.3 | Systems shall be suitable for examination of uncoated non-conductive materials and biological samples. |  |  |  |
| 1.1.0.2.4 | Systems shall be suitable for automatic analysis and examination of Gun Shot Residues (GSR). |  |  |  |
| 1.1.1 | Scanning Electron Microscope (SEM) Electron Gun and Optical System: |  |  |  |
| 1.1.1.1 | It shall have electron optical column system suitable for high resolution and high voltage studies. |  |  |  |
| 1.1.1.2 | The lenses in the column shall be air cooled. |  |  |  |
| 1.1.1.3 | Electron source shall be pre-centered Tungsten- (W) Hairpin Filament Type. |  |  |  |
| 1.1.1.4 | Acceleration voltage range shall be between 0.2 kV and 30 kV. |  |  |  |
| 1.1.1.5 | Probe current shall be at least 2 µA for Tungsten (W) Hairpin Filament Type. |  |  |  |
| 1.1.1.6 | Resolution shall be 3,0 nm or better at 30 kV at high vacuum for SE image and 4.5 nm or better at 30 kV at high vacuum and low vacuum for BSE image. |  |  |  |
| 1.1.1.7 | Magnification range shall be minimum 6x – 1000000x. |  |  |  |
| 1.1.2 | Sample chamber, sample stage and sample holders: |  |  |  |
| 1.1.2.1 | Diameter of sample chamber from left to right shall be at least 280 mm. |  |  |  |
| 1.1.2.2 | Systems shall have at least 8 ports for different detectors and accessories. |  |  |  |
| 1.1.2.3 | Sample stage shall be 5 axis (X, Y, Z, tilt, rotation), computer controlled, motorized and compucentric or eucentric geniometer type. |  |  |  |
| 1.1.2.4 | Rotationof the sample stage shall be 360° continuous. |  |  |  |
| 1.1.2.5 | Sample stage shall be able to move at least 100 mm in X and Y axis and 50 mm in Z axis. |  |  |  |
| 1.1.2.6 | Tilt angle of the sample stage shall be changed at least within 0° - (+)70°. |  |  |  |
| 1.1.2.7 | Sample holders shall be suitable for mono and multiple (at least 20) samples. |  |  |  |
| 1.1.3 | Vacuum system: |  |  |  |
| 1.1.3.1 | It shall be suitable for at least the following two different vacuum conditions: - High vacuum for examination of conductive and coated non-conductive materials (image and microanalysis) which shall be provided by a turbo molecular pump having capacity of at least 240 L/s and at least a rotary pump. - Low vacuum at least within 10 Pa – 2600 Pa. |  |  |  |
| 1.1.4 | Detectors: |  |  |  |
| 1.1.4.1 | The following image detectors and sample chamber observing device suitable for above mentioned vacuum system shall be given: - Secondary electron (SE) image detector for high vacuum. - A detector which works between 10-2600 Pa (low vacuum) in environmental mode must be given. - Back scatter electron (BSE) image detector suitable for both high and low vacuums. - A device suitable for observing sample chamber such as infrared camera (IR-CCD) or equivalent. |  |  |  |
| 1.1.5 | Energy Dispersive X-ray Spectrometer (EDS) |  |  |  |
| 1.1.5.1 | It shall be suitable to detect the elements between Be (4) and U (92), and for automatic GSR analysis. |  |  |  |
| 1.1.5.2 | Its detector shall be liquid nitrogen free Silicon Drift Detector (SDD) type and resolution shall be at least 126 eV or better (Mn Kα). |  |  |  |
| 1.1.5.3 | Its detector’s active area shall be at least 80 mm2 . |  |  |  |
| 1.1.5.4 | It shall be air cooled with Peltier cooling system. |  |  |  |
| 1.1.6 | System Computer A computer which is interoperable with the SEM-EDS system will be supplied. |  |  |  |
| 1.1.6.1 | Shall have Microsoft Windows 7 or 8 or 10 Professional Edition Turkish Version or equivalent |  |  |  |
| 1.1.6.2 | Shall have minimum 8 GB RAM |  |  |  |
| 1.1.6.3 | Shall have at least 1 GB graphics memory |  |  |  |
| 1.1.6.4 | Shall have internal DVD or Blu-ray optical drive |  |  |  |
| 1.1.6.5 | At least 128GB SSD and at least 500GB SATA |  |  |  |
| 1.1.6.6 | Shall have a minimum 21” LED or LED Backlight LCD monitor with at least 1600 x 900 pixel resolution, at least 5 ms response time. |  |  |  |
| 1.1.6.7 | Shall have at least 2 USB 3.0 ports, at least 2 USB 2.0 or higher ports, serial port, front headphone output, front microphone jack, DisplayPort (can be provided through an adapter), VGA video interface, Ethernet port. |  |  |  |
| 1.1.7 | Software |  |  |  |
| 1.1.7.1 | The licenses of all software shall be supplied. |  |  |  |
| 1.1.7.2 | Drivers of all of the external devices shall be supplied. |  |  |  |
| 1.1.7.3 | The following general analysis software shall be provided: - Quant ZAF and RhiRhio qualitative and quantitative analysis or equivalent - Quantitative analysis software without standard, - Radiation control software - SEM column control software - Scanning interface for SEM - Image collection and formation software - Additional software for quantitative mapping, line scanning, quick mapping and high resolution live spectral mapping, etc. |  |  |  |
| 1.1.7.4 | The Gun Shot Residue (GSR) analysis system shall be suitable to the above described system. |  |  |  |
| 1.1.7.5 | All necessary software and hardware for GSR analysis system to work with SEM shall be provided. |  |  |  |
| 1.1.7.6 | The GSR analysis software shall have at least the following properties: - Fully automatic GSR analysis and standard free quantitative analysis. - Automatic sample stage control - Three dimensional sample stage view for area selection. - Automatic X-ray mapping of GSR particles - Particle based X-ray mapping. - Spectral mapping with PTS technology. |  |  |  |
| 1.1.8 | Sample Coating Unit |  |  |  |
| 1.1.8.1 | It shall be suitable for at least two different types of coating (gold-palladium and carbon coatings). |  |  |  |
| 1.1.9 | Other Requirements |  |  |  |
| 1.1.9.1 | 1. 30 filaments. |  |  |  |
| 1.1.9.2 | 2. 100 conductive tapes |  |  |  |
| 1.1.9.3 | 3. 3 gold-palladium targets and 100m carbon fibers suitable for coating unit |  |  |  |
| 1.1.9.4 | 4. Necessary sample holders (1 holder with 1 sample capacity, 1 holder with 7-16 sample capacity, 1 holder with at least 20 sample capacity) must be given and necessary control software providing examination capabilities for various material sizes and types such as metals, plastics, powders, liquids, etc. including standard sample stages (Sb, Ba, Pb and standards related to GSR analysis). |  |  |  |
| 1.2 | Gas Chromatography/Mass Spectrometry System (GC/MS) The system will be used to make quantitative and qualitative analyses of psychotropic materials, toxic materials, explosives and unknown materials and will be bench-top type. |  |  |  |
| 1.2.0 | System Equipments: System consists of a gas chromatography, a mass spectrometer, an auto sampler, a personal computer and accessories. |  |  |  |
| 1.2.1 | Gas Chromatograph |  |  |  |
| 1.2.1.1 | Gas chromatograph shall be able to be controlled and programmed by computer. |  |  |  |
| 1.2.1.2 | It shall also be operable independent of computer. |  |  |  |
| 1.2.1.3 | It shall be designed for capillary column use. |  |  |  |
| 1.2.1.4 | Column oven shall be able to be heated up to at least 400 °C with 1 °C or lower increments, at least 15 steps temperature programs shall be applicable and it shall be capable of applying ramp rates between 1 and 100 °C/min. |  |  |  |
| 1.2.1.5 | Column oven must cool down from 400 °C to 50 °C at most within 5 minutes. |  |  |  |
| 1.2.2 | Mass Spectrometer |  |  |  |
| 1.2.2.1 | Mass scan range shall be at least 10-1050 amu. |  |  |  |
| 1.2.2.2 | The ionization shall be electron impact type ionization. |  |  |  |
| 1.2.2.3 | Ion source shall be able to be heated up to at least 300 °C. |  |  |  |
| 1.2.2.4 | It must have a quadrupole analyser. |  |  |  |
| 1.2.2.5 | Scan speed shall be at least 12.500 amu/sec. |  |  |  |
| 1.2.2.6 | Vacuum system with a turbomolecular pump with a capacity of at least 250 L/s. |  |  |  |
| 1.2.2.7 | Mass detector must be electron multiplier or photomultiplier type. |  |  |  |
| 1.2.2.8 | It must be sensitive for detecting 1 pg octafluoronaphtalene provided that S/N> 500/1. |  |  |  |
| 1.2.2.9 | There must be at least two filament installed in the system in EI mode. |  |  |  |
| 1.2.3 | Autosampler |  |  |  |
| 1.2.3.1 | Autosampler (standard sampler with at least 100 samples) shall be controllable and programmable with computer. |  |  |  |
| 1.2.3.2 | Vials which are used in the autosampler’s trays must have a volume between 0.5 ml - 2 ml. |  |  |  |
| 1.2.4 | System Computer A computer which is interoperable with the GC/MS system will be supplied. |  |  |  |
| 1.2.4.1 | Shall have Microsoft Windows 7 or 8 or 10 Professional Edition Turkish Version or equivalent |  |  |  |
| 1.2.4.2 | Shall have minimum 8 GB RAM |  |  |  |
| 1.2.4.3 | Shall have at least 1 GB graphics memory |  |  |  |
| 1.2.4.4 | Shall have internal DVD or Blu-ray optical drive |  |  |  |
| 1.2.4.5 | At least 128GB SSD and at least 500GB SATA |  |  |  |
| 1.2.4.6 | Shall have a minimum 21” LED or LED Backlight LCD monitor with at least 1600 x 900 pixel resolution, at least 5 ms response time. |  |  |  |
| 1.2.4.7 | Shall have at least 2 USB 3.0 ports, at least 2 USB 2.0 or higher ports, serial port, front headphone output, front microphone jack, DisplayPort (can be provided through an adapter), VGA video interface, Ethernet port. |  |  |  |
| 1.2.5 | Software |  |  |  |
| 1.2.5.1 | The original licence of the operating system shall be supplied. |  |  |  |
| 1.2.5.2 | Drivers of all of the external devices shall be supplied. |  |  |  |
| 1.2.5.3 | The software must perform complete instrument control of all GC and MS parameters, tuning, calibrations, diagnostic tests, data analysis during data acquisition, comparing ion chromatograms, common graphical representation of TIC and selected ions and mass ranges, searching mass spectra from the libraries, quantitative analysis, direct plotting of spectra and chromatograms, transfer of spectra and chromatogram to text editor software, etc. |  |  |  |
| 1.2.5.4 | Software shall include the latest and licenced versions of NIST STRUCTURES, WILEY, PFLEGER-MAURER-WEBER DRUGS commercial libraries; the library program must be able to create new libraries. |  |  |  |
| 1.2.6 | Other Requirements |  |  |  |
| 1.2.6.1 | 1. User level maintenance and service equipment kits, (column installation and cleaning kit, column cutting tools, ferrule picker, screwdriver, pliers etc.). |  |  |  |
| 1.2.6.2 | 2. Chemicals for calibration and controlling of instrument performance. |  |  |  |
| 1.2.6.3 | 3. Syringes for manual injection (1x5 µL, 5x10 µL) and for auto sampler (1x5 µL, 20x10 µL), septa (500 ea) resistant up to 300 °C and graphite ferrules (30) containing vespel, split/splitless “glass inserts” for injector or “glass liners” (30 items for each), enough quantities of screws, nuts, joints etc. for tie points, filaments for MS (20 items) and other equipment for ioniser, calibration and sensitivity materials etc., pump oil for pumps (at least 20 L for fore pump and if necessary 1 L for main vacuum pump). |  |  |  |
| 1.2.6.4 | 4. 20 ea 30 m x 0.25 mm ID, 0.25µm film thickness, fused silica, “low-bleed” type (suitable for MS), SE-54 phased (DB-5MS or equivalent) capillary columns. |  |  |  |
| 1.2.6.5 | 5. At least 8 m3, 2 Helium gas cylinders (99.999% pure) and 2 manometers suitable for the gas, filters, traps and fittings. |  |  |  |
| 1.2.6.6 | 6. At least 4 oxygen indicators, moisture, impurity and oxygen traps. 2 wall mounting kits for gas cylinders. |  |  |  |
| 1.2.6.7 | 7. 10000 vials, 10000 septa installed screw vial caps. |  |  |  |