CLARIFICATIONS TO TENDER DOSSIER NO: 1

Contract Title: 		Supply of Equipment for Strengthening Forensic Capacity of Turkey - Relaunch

Publication Reference: 	EuropeAid/137473/IH/SUP/TR

Q1]

Annex II-Technical Specifications

Item 1.1.5.2

“Its detector shall be liquid nitrogen free Silicon Drift Detector (SDD) type and resolution shall be at least 126 eV or better (Mn K)”

Please consider to change the above specification as below:

- “Its detector shall be liquid nitrogen free Silicon Drift Detector (SDD) type and resolution shall be at least 129 eV or better (Mn K)”

This specification is specific to one manufacturer and make impossible to propose other products other than that manufacturer. To be more competitive tender we kindly request you to change this specification.

A1]

The specification has been changed. Please refer to the “Changes to Tender Dossier No:1” document.

Q2]

Annex II-Technical Specifications

Item 1.1.5.3

“Its detector’s active area shall be at least 80 mm2”

Please consider to change the above specification as below:

- “Its detector’s active area shall be at least 60 mm2”

This specification is specific to one manufacturer and make impossible to propose other products other than that manufacturer. To be more competitive tender we kindly request you to change this specification.

A2]

Because of the required functionality, the specification remains without any change. In the market research, more than one product with eligible origin have been identified to satisfy the technical specifications.

Q3]

Annex II-Technical Specifications

Item 1.2.1.4

“Column oven shall be able to be heated up to at least 400 °C with 1 °C or lower increments, at least 15 steps temperature programs shall be applicable and it shall be capable of applying ramp rates between 1 and 100 °C/min.”

Please consider to change the above specification as below:

- “Column oven shall be able to heated up to 400 °C with 1 °C increments, at least 10 steps temperature programs shall be applicable and it shall be capable of applying ramp rates between 1 and 100 °C/min.”

Any method has not used over 3 steps as functionally. This specification make impossible to propose other products other than 2 specific products. By meeting the needs of end users and due to be more competitive tender, we kindly request you to change this specification.

A3]

Particularly in drug analysis, the temperature programs can be critical and analysts need 15 or higher steps temperature programs.

Because of the required functionality, the specification remains without any change.

Q4]

Annex II-Technical Specifications

Item 1.2.2.9

“There must be at least two filament installed in the system in EI mode.”

Please consider to change the above specification as below:

- “There must be at least two filament installed in the system in EI mode or if  the system has a single filament, It should be chanceable by user in less than 5  minutes.”

The device we want to offer has one filament. Please be informed that one filament is better than requested feature, the end user change the filament in less than 5 minutes. If the device has two or more filament only service can change them. By meeting the needs of end users and due to be more competitive tender, we kindly request you to change this specification.

A4]

In case there is only one filament, analysis will be interrupted when the filament is used up and it causes to stop the analysis for an intolerable time period which has a considerable impact on the effectiveness of the process.

Because of the required functionality, the specification remains without any change.

****